Hitachi, Ltd.

# Important Message

Before using this manual and the investigational device outlined in this manual, be sure to read A: Special Notes on page 41.

# For safe use

## Explanation of symbols

Precautions regarding safety are indicated by the following symbols and headings.

These precautionary signs consist of a safety alert symbol and the words “Warning” or “Caution.”

This is a safety alert symbol. It is used to warn people of potential danger of injury. In order to avoid injury or even death, follow the safety instructions below the symbol.

This is used to warn of potential danger that might lead to death or serious injury.

This is used to warn of potential danger that might lead to light or medium-level injuries.

This is used to warn of potential danger that might cause serious damage to the equipment or to surrounding property.

## General safety precautions

Read the following safety messages carefully and understand them fully.

- When operating the system, follow the instructions and procedures outlined in this manual.

- Always observe the precautions indicated on the equipment and in the manual.

Failure to observe these instructions or precautions may result in injury, fire, or damage to the equipment.

## Operation

Do not perform any operations or take any action not indicated in the manual.

If any problem occurs regarding the equipment, turn the power off, unplug the AC adapter, and contact your investigator.

## Precautions in handling and use

- Do not handle the AC adapter or plug it into an outlet with wet hands, as this may cause electrical shock.

- Use a power source from 100V AC to 240V AC (50-60Hz). Use of a voltage outside this range may cause smoke emission, fire, electrical shock, or failure.

- The user must never disassemble, modify, or repair the hardware.  Doing so may cause electrical shock, smoke emission, fire, or failure.

- The signal processing unit must always be used with the provided AC adapter (UE15WCP-050200SAP). Use of a different adapter may cause smoke emission, fire, electrical shock, or failure.

- Conditions such as those indicated below may lead to fire or electrical shock. Turn off the AC power source, then after confirming that the device is not too hot to touch, turn off the power switch, unplug the AC adapter from the outlet, and request that your investigator have the device repaired.

  - When smoke is emitted by the signal processing unit, LED/PD unit, or AC adapter, or when a strange odor is noticed.

  - When a foreign object gets inside the signal processing unit, LED/PD unit, or AC adapter, or when water,or other liquid is spilled on any of these.

  - When the signal processing unit, LED/PD unit, or AC adapter becomes abnormally hot.

- The amount of heat emitted by the LED is normally a safe level of 1(C or below. However, a problem in the signal processing unit, etc., may cause the temperature to rise slightly. Please observe the following precautions to prevent low-temperature burns.

  - Do not use while the patient is running a fever.

  - Avoid use in high room temperatures..

  - Do not wear the headband over inflamed skin.

  - Avoid prolonged use. Limit sessions to approximately 60 minutes each.

- The near-infrared light emitted and dispersed by the LED is a safe Class 1.  Nonetheless, observe the following precautions to reduce exposure to eyes and skin.

  - Stop use if skin irritation or rash, etc., is experienced in the area where the headband is worn.

  - Avoid looking directly at the red light emitted by the light source of the LED/PD unit.

    - The signal processing unit, LED/PD unit, and AC adapter contain precision electronic components. Avoid dropping or strong impact, which may cause damage or failure.

    - When unplugging the AC adapter from a power outlet, always grasp the adapter itself. Pulling on the cable may cause wire breakage and failure.

    - Always unplug the AC adapter of the signal processing unit from the power outlet when use is discontinued for an extensive period. While the AC adapter is connected to the power outlet, it continues to function as an AC adapter even when the controller is not in use, leading to wasteful energy consumption.

    - When connecting cables to the signal processing unit, whether from the LED/PD unit, the USB cable for connection to a computer, or the DC jack from the AC adapter, hold down the unit securely and make sure the connectors are pushed firmly all the way in. If the connectors are not pushed all the way in, unstable operation may occur.

    - Do not use excessive force that causes the LED/PD unit to bend, as this might cause damage to the internal circuit.

    - Using a soft cloth, periodically clean the filter on the LED/PD unit headband.

    - An inverter-type fluorescent lamp in the vicinity may interfere with the measured signals.  If interference occurs, turn the lamp off during use of the Kokoro Gatari, or move the lamp away from the Kokoro Gatari.

The computer to which the signal processing unit hardware is connected is to be used only for Kokoro Gatari. Avoid the following uses, which can cause faulty operation.

- Installing other software

- Connecting to the Internet or to a LAN without the investigator’s permission.

An isolated transformer of UL60601 acquisition should be set between the AC power supply and the computer for safety. The computer linked to Kokoro Gatari should be a product that is compatible with UL60950.

- Kokori Gatari is classified as Type B for electric shock according to UL60601-1.

- Disposal of the Kokoro Gatari does not present an environmental risk.  Please follow local ordinances or regulations for disposal of the Kokoro Gtari.

# CONTENTS

Important Message	1

For safe use	2

Explanation of symbols	2

General safety precautions	2

Operation	2

User’s precautions	2

Precautions in handling and use	2

CONTENTS	2

1. Overview	2

About Kokoro Gatari	2

Basic principles	2

Kokoro Gatari components	2

Measurement intervals and brain activation	2

2. Preparations	2

Cable connections and names of parts	2

Launching Kokoro Gatari	2

Patient name registration and selection	2

Shutting down Kokoro Gatari	2

3. “YES”/ “NO” Judgment	2

Data Gathering on the first day	2

Measure brain blood volume change	2

See the sample data positions on the Map	2

Display the Boundary and move it	2

Register the boundary	2

Acquire sample data for judgment by asking controlled questions	2

Data gathering from the second day onwards	2

“YES”/ “NO” boundary creation 	2

“YES”/ “NO” boundary registration	2

Additional sample data gathering	2

4. Additional Explanation	2

Measurement screen	2

“YES”/ “NO”-Boundary screen	2

History-Boundary and accuracy screen	2

Analysis screen	2

A: Special Notes	2

B: Troubleshooting	2

C: Hardware Specifications	2

# 1. Overview

## About Kokoro Gatari

Kokoro Gatari is a system for use with patients who have normal brain functioning but cannot move their body, and are therefore in what is called “totally locked-in state” (TLS). Its purpose is to obtain “Yes” and “No” responses from the patient to questions posed by family members or care-givers. The patient gives a “Yes” response by stepping up brain activity, and a “No” response by putting the brain into a relaxed state.

## Basic principles

“Yes” and “No” responses are judged by measuring blood volume at the frontal cortex, using near-infrared light. The near-infrared light produced by the light source is directed at the brain through the skull and scattered at the brain surface. Some of this light is reflected back through the skull, and is measured by the equipment.

## Kokoro Gatari components

The Kokoro Gatari system consists of a computer with software installed for judging “Yes”/ “No” responses and a controller (signal processing unit and a headband-mounted sensor unit). The sensor unit includes both a light-emitting diode (LED) for generating the near-infrared light and a photodiode (PD) for detecting returned light. The headband on which the sensor unit is mounted is worn so that the sensor contacts the skin at a location approximately 2 to 3 cm above the left eyebrow.

## Measurement intervals and brain activation

Kokoro Gatari obtains a single response by measuring changes in blood volume (returning light amount) over three intervals, which are referred to as a rest interval, response interval, and rest interval. Each interval lasts 12 seconds, so a measurement sequence of 36 seconds is required.

Rest interval 1: Prepare for the measurement

Response interval: Obtain a “Yes” or “No” response

Rest interval 2: Establish a clear distinction from the data obtained during the response interval

When a “Yes” response is to be given, the patient during rest interval 1 attempts to keep the blood volume steady by imposing a light load on the brain, such as by counting slowly. During the response interval, the patient activates the brain through a more intense activity such as performing a calculation (e.g., subtracting 3 from 100 successively) or by singing a song rapidly. The activity in rest interval 2 is a repeat of rest interval 1. By exercising the brain in this way, a wave pattern like that in the figure on the right is obtained.

To answer “No,” the patient maintains the same state throughout the three intervals, by imposing a light load such as counting slowly or slowly singing a children's song. By relaxing the brain in this way, a wave pattern like that in the figure on the right is obtained.

While the ideal state would be one free of all distracting thoughts, in reality a patient is likely to experience stray thoughts that cause the blood volume to increase. Achieving a state of stable blood volume by means of a light mental task of some kind is therefore preferable for obtaining satisfactory results.

# 2. Preparations

## Cable connections and names of parts

(1) Connect the sensor to the signal processing unit.

(2) Connect the signal processing unit to the computer by USB cable.

(3) Connect the provided AC adapter to the signal processing unit.

(4) Plug the AC adapter into an outlet from 100V AC to 240V AC.

Signal processing unit

Front view

Rear view

## Launching Kokoro Gatari

The procedures from powering up the computer to attaching the sensor to the patient’s forehead are described below.

(1) Push the power switch on the signal processing unit.

The signal processing unit power turns on.

Confirm that the power indicator (green lamp) on the signal processing unit is lit.

(2) Turn on the computer power. (For the location of the power switch, see the computer manual.)

Confirm that the USB indicator (green lamp) on the signal processing unit is lit.

In a few moments, the Kokoro Gatari main screen should appear.

(3) Note the contents of the "Choose patient name" field. A different patient name is set for each patient.

If no patient name is registered

Register the patient's name and choose that name (see page 2). The first time Kokoro Gatari is run, "user1" is set in the "Choose patient name" field.

When a patient name is already registered

Confirm that the patient name to be used is displayed in the "Choose patient name" field.* If it is not displayed, choose the correct name (see page 2).

＊：The name displayed initially in the "Choose patient name" field is the patient name registered when Kokoro Gatari was last shut down.

(4) Use the headband to position the center of the sensor unit approximately 2 to 3 cm above the left eyebrow of the patient.

After the headband is in place, confirm that the level indicator (green lamp) is lit. If it is not lit, tighten the headband so the sensor is firmly in contact with the skin.

Remember the sensor position so that the next time the headband is set the sensor is as close as possible to the same place.

### Patient name registration and selection

Patient name registration

(1) Enter a patient name.

(2) Click the “Register” button.

“Do you want to register patient name ‘Jack’?” is displayed.

Click the “Yes” button.

The patient name is displayed in the “Choose patient name” field.

Patient name selection (Usually the investigator registers the date/month/year as your ID for the study.)

(1) Choose a patient name from the drop-down text box (using ▼).

A list of patient names is displayed.

(2) Click on a patient name.

The drop-down list closes, and the selected patient name appears in the "Choose patient name" field.

## Shutting down Kokoro Gatari

(1) Click the “Quit” button in the main screen.

The following confirmation message appears.

(2) Click the “Yes” button to shut down.

The “Yes”/ “No” decision software shuts down, returning to the desktop display.

To run Kokoro Gatari again, double-click the Kokoro Gatari program icon.

(3) Click the “Start” button, then “Turn Off Computer.”

A Shutdown options dialog appears.

(4) Click the “Turn Off” button.

The computer power is turned off.

(5) Turn off the power on the signal processing unit.

The power indicator lamp is extinguished.

# 3. “YES”/ “NO” Judgment

In general, changes in the volume of patients’ brain blood differ from day to day. Sample data (i.e., responses to questions with known answers) to establish accuracy should therefore be gathered every day before judging “Yes”/ “No” answers.

Data gathering on the first day

Data is gathered on the first day according to the following 6 steps

(1) “YES”/ “NO” sample data gathering

Request known “Yes” responses two times and known “No” responses two times from the patient. The acquired data are referred to as “YES sample data” and “NO sample data,” respectively.

(2) “YES”/ “NO” sample data position

See the sample data positions on the Sample data distribution map (“the Map”).

(3) “YES”/ “NO” boundary creation

The “YES”/ “NO” boundary is necessary to judge “YES”/ “NO” sample data. The boundary on the Map displayed on the computer screen is manually drawn and should be moved so that the “YES” sample data are located on one side of the boundary and the “NO” sample data are located on another side as far apart and distinct as possible.

(4) “YES”/ “NO” boundary registration

Register the “YES”/ “NO” boundary.

(5) Controlled questions

Ask controlled questions that can be clearly answered “Yes” or “No” by the patient without hesitation. Judgment is carried out twice for “Yes” and “No” questions respectively.

(6) More sample data gathering

Only if a 70% accuracy rate is achieved, ask another 6 controlled questions and judge them with the device.

Data gathering from second day onwards

Data gathering from the second day is almost the same as on the first day, except for the following:

[Difference 1] “YES”/ “NO” sample data gathering

When “YES”/ “NO” sample data have been measured, the result of the “YES”/ “NO” judgment is displayed and heard on the device because the “YES”/ “NO” boundary already has been registered.

[Difference 2] “YES”/ “NO” boundary creation

The line that was selected from among the registered lines on the previous day is automatically shown on the Map (See page 2).

[Difference 3] “YES”/ “NO” boundary registration

When a new boundary is registered, the boundaries from previous days are shown with their accuracy rate on the History-Boundary and accuracy screen. If boundaries from a previous day are more accurate than the new one, the most accurate boundary is selected to judge answers (See page 2).

[Difference 4] Additional sample data gathering

If the accuracy rate is under 70%, move the “YES”/ “NO” boundary again so that it can separate “YES”/ “NO” sample data more correctly (See page 2). If the new boundary judges the answers more than 70% correctly, register the boundary and judge 6 more controlled questions using that boundary.

## Data Gathering on the first day

### Measure brain blood volume change

[Step 1] Display the Measurement screen and check the signal.

- Click the “Measure” button in the main screen.

The Measurement screen appears.

(2) Click the “Check signal” button

The software performs a preconfiguring operation.

When the signal check is finished, the results are shown on the screen. Look at the Check input signal on the screen and if “Possible” appears in the Data acquisition proceed to Step 2. If “Impossible” appears reset the sensor on the forehead.

[Step 2] Acquire “YES”/ “NO” sample data such as “YES”, “YES” , “NO”, “NO”

(i.e., a total of 4 measurements).

When acquiring “YES” sample data

(1) Set the Data mode to “YES” sample mode, if it has been set otherwise.

(2) After asking a “Yes” question to the patient, click the “Acquire data” button.

For making the patient answer “Yes”, see page 2.

Preconfiguring begins. Wait for 10 seconds until it ends.

When preconfiguration is complete, audible notification (a ringing tone) is given and data acquisition begins. At the transition from rest interval 1 to the response interval, a single tone is sounded; when going to rest interval 2, the same tone is sounded twice; and when data acquisition ends, the ringing tone is again heard. The number of “YES” samples is incremented by one when the sequence ends.

(3) Click the “End” button.

Clicking the “End” button causes the wave pattern to disappear, and the “End” button becomes the “Acquire data” button, enabling the next sequence to be started.

When acquiring “NO” sample data

(1) Set the “Data mode” to “NO” sample” if it has been set to the other modes.

(2) After asking a “No” question to the patient, click the “Acquire data” button.

For making the patient answer “No,” see page 2.

The following steps are the same as when acquiring “YES” sample data.

### See the sample data positions on the Map

- Click the “Set boundary” button on the Measurement screen.

The “YES”/ “NO” boundary screen is displayed.

Four sample data values are shown on the Map.

indicates “YES” sample data. The number on the right is the data number*.

indicates “NO” sample data. The number on the left is the data number*.

indicates the median point of the “YES” sample data.

indicates the median point of the “NO” sample data.

＊：Data numbers correspond with the numbers in the “YES”/ “NO” sample data table.

Sample data file name

The first 14 digits indicate the date and time of data acquisition. For example, the file name:

20081001091414Y

indicates that the data were acquired on October 1, 2008 at 9:14:14 AM. Y and N indicate “YES” and “NO” sample data, respectively.

### Display the Boundary and move it

(1) Click the “Manual boundary” check box.

A straight boundary is shown on the Map.

A boundary divides the plain on the Map into two areas:

the area with ■ is called the “YES” area and the area with ■ is called the “NO” area.

On the above Map, the left side of the boundary is the “YES” area and the right side is the “NO” area. Check that the “Yes” area has equal or more “YES” sample data than the “NO” area, and vice versa.

If the “YES” area has less “YES” sample data than the “NO” area and the “NO” area has less “NO” sample data than the “YES” area, click the Reverse “YES”/ “NO” button to switch the “YES” and “NO” areas.

- Move the boundary so that the “YES” area has only “YES” sample data and the “NO” area has only “NO” sample data, as far as possible.

①　Click the “Move” radio button to be able to move the boundary.

②　Point the mouse cursor at the boundary.

③　Drag & drop the boundary to a position so that the   1 is located in the “NO” area.

### Register the boundary

(1) Click the “Close” button.

A message appears asking whether you want to register the boundary.

(2) Click the “Yes” button.

The Measurement screen is displayed.

- Click the “See history” button.

The History-Boundary and accuracy screen is displayed.

- Click the “OK” button.

The Measurement screen is displayed again.

### Acquire sample data for judgment by asking controlled questions

[Step1] Ask 4 questions that can easily be answered “Yes” or “No” in a sequence such as “YES,” “NO,” “NO,” “YES,” and have them judged by the device.

(1) Confirm that “YES” sample mode is selected in the Data mode. If the other mode is selected click the “YES sample mode” radio button.

(2) Click the “Acquire data” button after asking a controlled “YES” question (e.g., Were you born in April?)

Data acquisition now starts. After data acquisition is completed, the results are given audibly and in the screen below.

(3) Click the “End” button.

(4) Click the “NO sample mode” radio button in order to change the mode from “YES sample mode.”

(5) Click the “Acquire data” button after asking a controlled “NO” question (e.g., Were you born in December?).

Data acquisition now starts. After data acquisition is completed, the results are given audibly and in the screen below.

(6) Click the “End” button.

Now the first data “YES” and the second data “NO” have been acquired. Continue acquiring the third data “NO” and the fourth data “YES” in the same way.

[Step2] Check that the accuracy rate is at least 70%.

(1) Click the “See history” button in the Measurement screen.

The History–boundary and accuracy screen is displayed.

Today’s accuracy rate is shown in the Today % column. It is calculated from all the data acquired on that day, which includes 4 data samples originally measured to create a boundary (called the “original data”) and 4 data samples judged (called the “test data”).

Origin % column shows the accuracy of the original data

(2) Click the “OK” button.

The Measurement screen appears again.

## Data gathering from the second day onwards

From the second day, data gathering follows almost the same procedure as on the first day. The differences mentioned on page 2 are explained in detail here.

### “YES”/ “NO” boundary creation

The boundary on the Map that was selected last from the registered ones on the previous day is displayed automatically. See the graph below.

### “YES”/ “NO” boundary registration

The History–Boundary and accuracy screen shows the accuracy rate of all registered boundaries, including not only the boundary today but also ones registered on previous days. After a boundary is registered, it is automatically selected to judge the test data and the selection is confirmed on this screen in the Selection column.

After a boundary is registered it is automatically selected, although it may not be used to judge test data.

The boundary with the highest accuracy rate should be selected manually as follows:

(1) Click on the boundary with the highest accuracy rate.

The above list shows that the boundary registered on the 4th day is automatically selected, not the boundaries created on the previous 3 days.

The boundary pointed by the mouse cursor turns blue to show it is selected.

(2) Click the “OK” button.

The Measurement screen appears again.

### Additional sample data gathering

After acquiring the test data a new boundary can be registered.

If the Today’s accuracy rate is under 70% try to create a new boundary. The currently selected boundary can be moved so that it separates the test data with a higher accuracy rate. If it is successful, register that new boundary (For how to register a boundary, see page 2).

(1) Click the “Set boundary” button in the Measurement screen.

“YES”/ “NO”-Boundary and accuracy screen is displayed.

In the above graph, the   3 (“YES” sample data) , the   3 (“NO” sample data) and the   4 are not correctly judged.

(2) Move the boundary ( For how to move a boundary, see page 2).

In this graph, the   3 and the   3 are not correctly judged.

The accuracy rate = 75%

(3) Rotate the boundary (For how to rotate a boundary, see page 2).

In this graph, all test data are judged correctly.

The accuracy rate = 100%

# 4. Additional Explanation

This chapter provides information about operations and functions that are not explained in previous chapters. Please review and use as a reference when necessary after becoming proficient in operating the system.

## Measurement screen

Data mode

- Decision mode

This mode can be used for actual “Yes”/ “No” response decisions. The decision mode should not be used in the study.

- Accuracy

Clicking this button displays the accuracy rate of today’s acquired sample data in the accuracy rate field.

## “YES”/ “NO”-Boundary screen

(1) Specify date

Enter the period of days within which sample data are displayed in the sample data distribution map (the Map) and the sample data table. For example, input 3 into the text box and click the “Update list” button to display sample data from today and the previous 3 days.

(2) Waveform display of sample data and changing data mode status

The waveform screen is displayed by double-clicking a sample data file name. The data status can be changed on this screen.

For data analysis, acquired data are classified into the following statuses: Sample, Used, Excluded, New, and Old. All data collected in the study are sample data.

(3) Display mode

Display mode is set to specify the status of sample data displayed on the Map. Only data corresponding to statuses selected in the Display mode are displayed on the Map.

(4) Line type

Parabola

Clicking this radio button displays a parabola on the Map.

Quadratic curve

If there are sample data classified as ‘Used,’ a quadratic curve automatically calculated from that ‘Used’ sample data is displayed. This option should not be used in the study.

(5) Edit

Rotation

Click the “Rotate” button and then move the mouse cursor to the boundary.

You can now drag the boundary to rotate it.

Bend

When a parabola or quadratic curve is chosen as the Line type, the curvature of the boundary can be changed in the same way as rotation. Move the mouse cursor to the boundary and drag it to bend it.

(6) Analyzed segments

Analyzed segments are time segments for analyzing waveform data for use in “Yes”/ “No” decisions. Analyzed segments are set by study investigators. Analyzed segments are defined by setting a start and end point. Those points are the elapsed time (sec) from the start of measurement. The initial values of the start point and end point are set to 15 sec and 35 sec respectively.

Auto calc.

The start points and end points are automatically set to show the maximum difference in the Y and N decision data waveform.

Update

After changing a start point and/or an end point manually or automatically, click this button to update the Map.

(7) Initial state

Click this button to return the “YES/NO boundary” screen to the initial state it was in when opened.

## History-Boundary and accuracy screen

① Comment

Display the text box by double-clicking a line in the boundary list. Comments on a boundary can be input into the text box

② Past accuracy

Click this button to display the “History-Accuracy” screen. On this screen, you can view the accuracy of past sample data based on a selected boundary.

③ Today’s T/F

Click this button to view today’s True/False results and errata for today’s sample data.

④ Delete

Click this button to delete a boundary selected in the table.

## Analysis screen

① Open

A window for choosing files is displayed.

Select the file name of the data to be viewed and then click the “Open” button.

The Analysis screen is displayed showing where the decision data fits on the Map, the decision data waveform, and the “Yes”/ “No” decision result.

Click the “Zoom” button to see an expanded view of the waveform.

# A: Special Notes

1. Limitations on use

- “Yes”/ “No” decisions by Kokoro Gatari are not guaranteed to be 100 percent accurate.

2. Suitability of patients/operators

- The Kokoro Gatari system is suitable for patients who meet the following conditions.

  - Patients who can hear care-givers’ questions and understand the meaning of the questions.

  - Patients with clear consciousness who can intentionally activate or de-activate the brain (put the brain in a relaxed state).

- Regardless of the above, there are certain patients with a low accuracy rate with whom the Kokoro Gatari system cannot be used. First take “YES”/ “NO” sample data and, based on the accuracy rate, decide whether or not the system can be used.

- The explanations in this manual are for those who can operate a Windows® computer. Those who cannot do so are requested to first read the computer manual.

Windows® is the registered trademark of U.S.-based Microsoft Corporation in the United States and other countries.

3. Guarantee

- In the event that the Kokoro Gatari system malfunctions during the guarantee period, repairs will be conducted in accordance with the “Kokoro Gatari Statement of Limited Warranty.” Please see this statement for details.

- For matters concerning the computer, please see the computer’s statement of guarantee.

4. Inquiries about this instruction manual

If you do not understand any instructions in this manual or if you find any errors in it, please contact us at the following e-mail address.

[E-mail]

aseg@ml.itg.hitachi.co.jp

[Office]

Business Incubation Division, Hitachi, Ltd.

6,6 Marunouchi 1-chome, Chiyoda-ku, Tokyo, 100-8280 Japan

Tel: +81-3-4564-2723  Fax: +81-3-4564-1444

5. Intellectual property rights

Reproduction or copying of part or all of this instruction manual without permission is prohibited.

6. Revision of text

The text of this instruction manual may be revised in the future without prior notice.

7. Rules/measures

Harmful interference

Note: This equipment has been tested and found to comply with the limits for a Class B digital device, pursuant to Part 15 of the FCC Rules. These limits are designed to provide reasonable protection against harmful interface in a residential installation. This equipment generates, uses, and can radiate radio frequency energy and, if not installed and used in accordance with the instructions, may cause harmful interference to radio communications. However, there is no guarantee that interference will not occur in a particular installation. If this equipment does cause harmful interference to radio or television reception, which can be determined by turning the equipment off and on, the user is encouraged to try to correct the interference by one or more of the following measures:

-- Reorient or relocate the receiving antenna.

-- Increase the separation between the equipment and receiver.

-- Connect the equipment into an outlet on a circuit different from that to which the receiver is not connected.

-- Consult the dealer or an experienced radio/TV technician for help.

8. UL60601-1 Compliance Statement

The Kokoro Gatari has been investigated by Underwriters laboratories for use within the patient environment. At the back of the Kokoro Gatari, it is shown that this unit is listed by the simple label notation at UL.

9. Contact number in event of malfunction

For the contact number in the event of a malfunction, please see the “Kokoro Gatari Statement of Limited Warranty.” If you are using the system on loan, please contact your investigator .

# B: Troubleshooting

NOTE: the “system manager” referred to below is your investigator.

Problem: The power indicator does not light up.

| Points to check | Countermeasures |
| --- | --- |
| 1. Is the power switch on the rear of the signal processing unit turned on? | Turn the power switch on. |
| 2. Is the electric power connector from the AC adapter firmly inserted in the DCIN +5V connector on the rear of the signal processing unit? | Insert the electric power connector firmly all the way in. |
| 3. Is the AC adapter firmly connected to the AC power outlet? | Firmly connect it to the power outlet. |
| 4. Is the AC power outlet supplied with power? | Ensure that the AC power outlet is supplied with power. |
| 5. The AC adapter may be malfunctioning. | Contact the system manager and have the AC adapter replaced. |
| 6. The signal processing unit may be malfunctioning. | Contact the system manager and have the unit replaced. |

Problem: The level indicator does not light up (the power indicator is lit).

| 1. Is the sensor (LED/PD unit) firmly inserted in the sensor connector on the front of the signal processing unit? | Insert the sensor connector firmly. |
| --- | --- |
| 2. Is the sensor attached at the place on the forehead indicated in the manual? | Check the place at which the sensor is attached to the forehead and the contact between the sensor and the skin. |
| 3. The sensor may be malfunctioning. | Contact the system manager and have the sensor replaced. |
| 4. The signal processing unit may be malfunctioning. | Contact the system manager and have the unit replaced. |

Problem: The USB indicator does not light up (the power indicator is lit).

| 1. Is the USB-connected computer’s power switch turned on? | Turn on the computer’s power switch. |
| --- | --- |
| 2. Is the supplied USB cable connector firmly inserted in both the signal processing unit and the computer? | Firmly insert the USB connector. |
| 3. The signal processing unit may be malfunctioning. | Contact the system manager and have the unit replaced. |
| 4. There may be a problem with the computer connected to the system. | Contact the system manager and have the computer replaced. |

Problem: Waveforms are not displayed by the computer (the power indicator and level indicator are lit)

| 1. If measured waveforms do not appear on the computer screen when all the LED on the panel are lit, select “Measure” from the initial menu screen to display the measurement screen. Check that “Input Signal Normal” is indicated in the green display area at the top right of the screen. | If the system still does not seem to be functioning normally even though “Input Signal Normal” is indicated, contact the system manager for guidance. |
| --- | --- |
| 2. “Outside Input Signal Range” is indicated in yellow in the above-mentioned green display area. | The system is not functioning normally. Contact the system manager for guidance. |
| 3. “No Input Signal” is indicated in red in the above-mentioned green display area. |  |

# C: Hardware Specifications

1. Components

The Kokoro Gatari hardware includes the following components.

(1)	Headband-mounted sensor (LED/PD unit)	HS-501A

(2)	Signal processing unit			YN-501A

(3)	Exclusive AC adapter 			UE15WCP-050200SPA

(4)	USB cable 				Cable with USB connectors on both ends

(A connector: male; B connector: male)

Note: A separate exclusive computer is required to use the hardware.

2. Main Specifications (Electrical Characteristics)

(1) Sensor unit light emitting diode (LED)

LED wavelength (2 wavelengths used): 770 (840) nm

LED output: approx. 10 mW (Class 1)

(2) Sensor unit photodiode (PD)

PD minimum incoming light: 10 PW (rms)

PD active area diameter: 1.2 mm

(3) Signal processing unit lock-in amplifier

Input amplifier gain: (20 (pins 3-4 jumpers)

(4) Signal processing unit AD converter

Full-scale input: 5Vp-p

Resolution: 16 bits

(5) Signal processing unit AC adapter

Model number: UE15WCP-050200SPA

Input : 100V AC to 240V AC 50-60Hz 400mA

Output voltage: +5V±5%, 2A

This AC adapter is ClassⅡ

Use in combustible gas atmosphere: It is not suitable

(6) Other specifications

Operating temperature range: 5(C to 35(C

Operating humidity range: 35％ to 80％（No Condensation）

In transit/stock temperature range: 0℃ to +45℃

In transit/stock humidity range: 10％ to 85％（No Condensation）

In transit/stock Atmosphere range: 900hPa to 1080hPa

Power consumption: approx. 5.0V DC 0.8A

Dimensions: 198 m (W) x 39 mm (H) x 181 mm (D)

Weight: approx. 1.0 kg

Kokoro Gatari Instruction Manual for Study

1st edition October 2008

2nd edition January 2009

Business Incubation Division, Hitachi, Ltd.

6,6 Marunouchi 1-chome, Chiyoda-ku, Tokyo, 100-8280 Japan

# Text boxes

System for detecting “Yes”/ “No” responses based on cerebral blood volume changes

Kokoro Gatari

(Type: YN-501A)

Instruction Manual

for Study

CAUTION: Kokoro Gatari is limited by U.S. law to investigational use.

#### WARNING

#### CAUTIONNG

#### CAUTIONNG

#### WARNING

#### CAUTION

#### CAUTION

##### Reflected light

(Output light)

LED

Photodiode

Near-infrared light 
(Input light)

Skull

Active site (blood volume changes)

Cerebral 
Cortex

Headband-mounted sensor (see page 9)

Computer with “Yes”/ “No” judging program installed

Signal processing unit

12s

12s

12s

Rest 2

Rest 1

Response

Sensor

(LED/PD unit)

Plugs into sensor connector on front of signal processing unit

(LED/PD)

Velcro tape

Velcro tape on reverse side

To USB port on COMPUTER

(For the location of the USB port, refer to the COMPUTER manual.)

Plugs into USB cable connector on rear of signal processing unit

(USB OUT)

Plugs into AC adapter connector on rear of signal processing unit

(DC IN +5V)

AC adapter

Sensor connector

Power indicator

Level indicator

Part of ME EQUIPMENT that in NORMAL USE comes into physical contact with the PATIENT to perform its function.

USB indicator

Power switch

USB cable connector

Note: The GND terminal (screw) does not need to be connected.

AC adapter connector

Program icon

Note

When acquiring data for the first time after connecting the sensor unit, click the “Check signal” button. It takes approximately 10 seconds for the “Check signal” button to become active due to the initialization of the device.

Results

Note

Only perform this signal check once, before the first data measurement.

When this item is clicked, “YES” sample mode is selected.

Note

The first time you start up the Kokoro Gatari software, “YES” sample is selected by default.

Note

Preconfiguring is carried out in order to check the intensity level of the reflected light.

12s秒

12s秒

Response

Rest 2

12s秒

Rest 1

Incremented by 1

Note

The “End” button can be used in case a loud noise or other interference occurs during data acquisition, causing a sudden change in the waveform.

When this item is clicked, “NO” sample mode is selected.

Note

The mode selected in the previous operation is automatically selected.

“YES”/ “NO” sample data table

Indicates “NO” area.

###### Boundary

Indicates “YES” area.

Note

A straight line is initially selected as the default.

■ and ■ have been switched

Note

A boundary can be operated as follows:

① It can be rotated (See page 42).

② It can be drawn as a parabola (See page 42).

If the first row is displayed, the boundary has been registered

Note

If you want to ask additional questions for which the response are not known that can only be answered by the patient, click the “Decision mode” radio button. Decision mode will not be used in this study.

Result

Result

Note

By clicking the “Today’s T/F” button, you can view all that day’s sample data and True/False judgments.

Original data

Test data

Boundary on first day

Boundary on next day

Boundaries are displayed in order of newness, (with the newest at the top）

“YES” test data

“NO” test data

①

②

(1)

(3)

(2)

(4)

(5)

(7)

(6)

Note

① Data status is set in the sample data list. The status of data acquired in a study is sample data.

② To know the date and time of a boundary, close the “YES”/ “NO” boundary screen to return to the Measurement screen, and then click the “See history” button.

Acquired data (blue)

Data after low-pass filtering with cutoff frequency of 0.1Hz (pink)

Baseline (green)

End

Start

Rest 1

Response

Rest 2

Analyzed

segment

Start of measurement

③

④

②

①

①

Decision data file names. Unlike with sample data files, Y or N is not appended after the first 14 characters.

Note

In the study, nothing is displayed in the Data plotting area and the “YES”/ “NO” boundary and model data distribution area.

MEDICAL　EQUIPMENT

WITH　RESPECT　TO　ELECTRICAL　SHOCK,

FIRE　AND　MECHANICAL　HAZARDS　ONLY

IN　ACCORDANCE　WITH　UL60601-1 3NTZ

All rights reserved